[image: image70.jpg]%VKK?



®

Gubbs Mass Spec Utilities / Quick Calc™ 6.x.x

Administration and Installation Manual
Page 5 of 83

	Gubbs Mass Spec Utilities / 

Quick Calc™ 6.x.x
GMSU/QC 6 and earlier

	Supported Platforms


Thermo® XDKOEM 1.0.2.15


Sciex® Analyst™ 1.3.1 - 1.4.2


	Administration and Installation Manual

	Gubbs, Inc.

265 Blue Spruce Circle
Alpharetta, GA 30005
P: 770-573-0169
F: 508-453-1338



	gubbs@gubbsinc.com
www.gubbsinc.com




	Table of Contents

Section
Page Number


	21
Introduction


22
Definitions


23
High Throughput (HT) vs. Low Throughput (LT) Acquisition


24
Data File Convention


25
Data Acquisition Method Assumptions


26
Multiple Analyte Experiments


27
Data Acquisition Data File Content Rules


27.1
Sciex


27.2
Thermo


28
Thermo Processing Methods


29
Installation Overview


29.1
Inspect system requirements and prerequisites


29.2
Execute Setup.exe


29.3
Determine the physical location of the GMSU/QC database Gubbs_01.mdb (if applicable)


29.4
Prepare Windows® directory share and directory permissions (if needed)


29.5
Configure database (if needed)


29.6
Configure DNS ODBC file (if needed)


29.7
Execute UpdateGMSU/QCDatabase.xls (UGD) (if applicable)


29.8
Configure GMSU/QC Configuration Default Parameters


210
System Requirements


211
Pre-Requisites


211.1
Sciex environment


211.2
Thermo Scientific environment


212
Setup Execution


212.1
Setup Execution


212.2
End User License Agreement (EULA)


212.3
Query for database type


212.4
Installation Finish


212.5
Additional Actions Performed by Setup


213
GMSU/QC Licensing


213.1
Licensing Process


214
Physical Location of GMSU/QC database Gubbs_01.mdb


214.1
Stand-alone Environment


214.2
Multi-user Environment


215
Windows Share Permissions and Directory Permissions


215.1
Shared Network Data Directory (if applicable)


215.2
Shared Local Data Directory (if applicable)


216
Database Configuration – Microsoft® Access


217
DSN Configuration – Microsoft® Access


217.1
ODBC Data Source Administrator


217.2
Create New Data Source


217.3
ODBC Setup


217.4
Select database


218
Database Configuration – Microsoft® SQL Server


218.1
Create a SQL Server database using the MS Access Upsizing Wizard


218.2
Create a GMSU/QC database account


218.3
Record the user account and password information


218.4
Local Data Source Name (DSN)


219
DSN Configuration – Microsoft® SQL Server


219.1
ODBC Data Source Administrator


219.2
Create new data source


219.3
Create a new data source


219.4
Enter connection parameters


219.5
Additional parameters


219.6
Test connection


220
GMSU/QC_xx.ini


220.1
MS Access GMSU/QC_xx.ini configuration


220.2
SQL Server GMSU/QC_xx.ini configuration


221
Gubbs Inc Password Encryption Tool


222
UpdateGMSU/QCDatabase.xls (UGD)


222.1
Description of UGD contents


222.2
Procedure


223
Configuring Test and Production GMSU/QC Database Instances


223.1
Procedure


224
GMSU/QC Configuration Default Parameters


225
Console


226
Configuration Utility


226.1
Caco - Global


226.2
Caco – Analysis Time Points


226.3
Caco – Chromatography Peak Sets


226.4
Caco – Constants


226.5
Flags


226.6
Generic Viewer


226.7
Global


226.8
Hepatic Clearance Calculator


226.9
Hepatic Clearance Calc. Expt. Constants – Existing Experiments Tab


226.10
Hepatic Clearance Calc. Expt. Constants – Add Species/Type Tab


226.11
Hepatic Clearance Calc. Expt. Constants – Experiment Time Points Tab


226.12
Modify Method


226.13
PK Calculator - Global


226.14
PK Calculator – Species


227
Database Table Structures and Definitions


228
References






	Table of Appendices
Appendix
Page Number


	2Appendix A
Thermo Environment Multiuser System Configuration






	Table of Figures
Figure
Page Number


	2Figure 1
License Prompt


2Figure 2
MS Access DNS - Data Source Administrator


2Figure 3
MS Access DNS - Create a New Data Source Wizard


2Figure 4
MS Access DNS - ODBC Microsoft Access Setup


2Figure 5
MS Access DNS - Select Database


2Figure 6
MS Access Upsizing Wizard


2Figure 7
MS Access Upsizing Wizard - Choose Create new database


2Figure 8
MS Access Upsizing Wizard - Connection settings


2Figure 9
MS Access Upsizing Wizard - Export all tables


2Figure 10
MS Access Upsizing Wizard – Export attributes


2Figure 11
MS Access Upsizing Wizard – No application changes


2Figure 12
MS Access Upsizing Wizard - Finish


2Figure 13
MS SQL Server GMSU/QC account roles


2Figure 14
SQL Server DNS - Data Source Administrator


2Figure 15
SQL Server DNS - Create a New Data Source Wizard


2Figure 16
SQL Server DNS - Create a new data source


2Figure 17
SQL Server DNS - Enter connection parameters


2Figure 18
SQL Server DNS - Enter additional parameters


2Figure 19
SQL Server DNS - Enter additional parameters


2Figure 20
SQL Server DNS - Test connection


2Figure 21
SQL Server DNS - Test connection


2Figure 22
GMSU/QC UGD notification


2Figure 23
UGD Worksheet ‘Sheet 1’


2Figure 24
UGD Worksheet ‘ini_Path’


2Figure 25
UGD – After Execute Update


2Figure 26
UGD – SQL Statements


2Figure 27
GMSU/QC Console


2Figure 28
Configuration Utility – Caco (Permeability Calculator) – Global


2Figure 29
Configuration Utility – Caco (Permeability Calculator) – Analysis Time Points


2Figure 30
Configuration Utility – Caco (Permeability Calculator) – Chromatography Peak Sets (shown in Edit mode)


2Figure 31
Configuration Utility – Caco (Permeability Calculator) – Constants


2Figure 32
Configuration Utility - Flags


2Figure 33
Configuration Utility – Generic Viewer


2Figure 34
Configuration Utility - Global


2Figure 35
Configuration Utility – Hepatic Clearance Calculator


2Figure 36
Configuration Utility – Hepatic Clearance Calc. Expt. Constants – Existing Experiments Tab


2Figure 37
Configuration Utility – Hepatic Clearance Calc. Expt. Constants – Add Species/Type Tab


2Figure 38
Configuration Utility – Hepatic Clearance Calc. Expt. Constants – Experiment Time Points Tab - Initial


2Figure 39
Configuration Utility – Hepatic Clearance Calc. Expt. Constants – Experiment Time Points Tab – Edit Mode


2Figure 40
Configuration Utility – Modify Method


2Figure 41
Configuration Utility – PK Calculator – Global


2Figure 42
Configuration Utility – PK Calculator– Species






1 Introduction

Gubbs Mass Spec Utilities / Quick Calc (GMSU/QC) is a suite of utilities that enhance the ability of scientists to perform post-acquisition processing of data acquired in a high throughput (HT) environment.
GMSU/QC uses either Microsoft® (MS) Access or MS SQL Server as a datastore. MS Access may be used to quickly setup a GMSU/QC instance for user testing or production use. MS SQL Server may be used if a more secure and possibly more efficient environment is desired. The MS Access database can be directly upgraded to SQL Server without loss of data. If users desire custom information reports, data stored in the database is easily accessible.
To perform chromatographic functions, GMSU/QC uses the software development kit components provided by Sciex and Thermo, respectively. Therefore, when users optimize chromatography, they can expect the features available in the native chromatography data system (CDS) will be behave similarly in GMSU/QC. In the same manner, it can be expected that the GMSU/QC chromatographic features will differ between Sciex data and Thermo data, though every effort has been made to ensure the differences are minimal.

GMSU/QC may be used in a multi-user environment. There are several configuration possibilities for multi-user environment which are described in more detail in Section 15. Briefly, the GMSU/QC database and acquired mass spectrometer raw data files may be stored in a common location accessible to others (either a shared file server directory or a shared local directory). In both instances, it is most optimal to place the GMSU/QC Access database on a filer server accessible to all GMSU/QC users.

In a shared file server directory configuration, multiple installations of GMSU/QC may be configured to use the database and access the raw data files located in the common location. This allows users in the analytical department to use GMSU/QC to optimize chromatographic integration from their office rather than the laboratory. Once the chromatographic integration data is approved, users in the pharmacokinetics department may be notified and can immediately generate pharmacokinetic data.
In a shared local directory, GMSU/QC may be configured to “Use UNC Path” (alternatively admins may map each shared local directory to a unique letter). With this setting, data may remain at the local data acquisition workstation and all data pointers in the GMSU/QC database will reference the complete UNC directory path, rather than “C:\”, for example. Please note that, in this configuration, the local data acquisition data directory (e.g. …\Analyst Data\, or …\XCalibur\) must be “shared” to all GMSU/QC users. An Information Technology (IT) representative can assist in this configuration (see Section 15.2 for more details). Also, see Appendix A
Thermo Environment Multiuser System Configuration, for a detailed discussion on configuring GMSU/QC in a multiuser environment.
The most current version of this manual may be found by browsing to www.gubbsinc.com and navigating to the GMSU/QC page.

2 Definitions

	Term
	Definition

	CDS
	Chromatography Data System

	DSN
	Data Source Name

	EULA
	End User License Agreement

	GMSU/QC
	Gubbs Mass Spec Utilities

	HT
	High Throughput

	LT
	Low Throughput

	MS
	Microsoft®

	UGD
	UpdateGMSU/QCDatabase.xls


3 High Throughput (HT) vs. Low Throughput (LT) Acquisition

One of the salient features of GMSU/QC is that it can process HT acquired data as well as LT acquired data.

Typically, mass spectrometry data acquisition systems require that compound(s) of interest occur once in a chromatographic analysis. Therefore, in order to analyze a 12-point assay, scientists had to acquire twelve separate ‘samples’. This is referred to as LT acquisition in the GMSU/QC system.

In HT acquisition, users open the chromatography data system (CDS) data acquisition window for as long as needed and inject and acquire all assay time points in a single chromatographic run. The benefits of HT acquisition 1 - 4 include:

· The ability to view all chromatographic peaks simultaneously to check for possible instrumental or sample handling problems and/or trends

· When combined with multiple-column switching techniques, increases throughput by decreasing the time between sample injections and minimizes communication problems that can occur between instrument and data acquisition system.

NOTE: Currently HT acquisition data processing is supported on the Sciex platform in the Hepatic Clearance and Permeability modules only.

It must be noted here that half of the benefit of HT acquisition is addressed by the ‘Display HT Chromatography’ feature of GMSU/QC. If normal LT acquisition is performed, users may click on the ‘Display HT Chromatography’ button to view all the injections (or samples) of an assay in a single chromatogram. GMSU/QC retrieves the XY data of each chromatographic injection/sample and concatenates the data end-to-end in a text file. Both ABI and Thermo chromatographic display components used by GMSU/QC allow text data to be loaded and viewed as a chromatogram. See the GMSU/QC Users Manual – Display HT Chromatography section for more information.

4 Data File Convention

Throughout this manual, the term “Data File” is used.

In the Sciex environment, the Data File is the .wiff file.

In the Thermo environment, the Data File is the .sld file. Even though the .sld file doesn’t actually contain raw data (.raw files), the .sld contains the directory path reference to the .raw files. Since the .sld file “manages” the location of the .raw files, the .sld will be referred to as the Data File.

5 Data Acquisition Method Assumptions

GMSU/QC makes the following assumption concerning the configuration of Analyte and Internal Standard transition ions within the data acquisition method (applies to single-analyte assays):

The Analyte is listed first in the experiment transition ion section.

The Internal Standard is listed second in the experiment transition ion section.

If the opposite is true or the Chromatography Data System (CDS) lists the transition ions in order of compound molecular weight, then the user has the ability to assign the appropriate transition ions to Analyte and Internal Standard using the Menu item ‘Assign Transition Ions’ located in each module.

6 Multiple Analyte Experiments

The ‘Assign Transition Ions’ feature can also be used to prepare data in which data for multiple analytes have been acquired (e.g. in cassette dosing experiments or looped experiments).

Please note that this process is automated in the Thermo environment in which a valid processing method (.pmd) has been configured in the sequence file (.sld). The .pmd can be created manually using Thermo Xcalibur™, automatically using Thermo QuickQuan™, or exported with the .sld from Thermo LCQuan™. Please see the relevant help files inside any of these CDS’s for further information.

7 Data Acquisition Data File Content Rules

Data file content must conform to certain rules in order to be successfully processed by GMSU/QC. The rules differ slightly depending on the data file.

7.1 Sciex

· Sciex data must be stored in a single .wiff file

· The .wiff file may contain data only from a single assay

· An ‘assay’ is defined as the injections/samples that make up the data set for that analysis (e.g. Hepatic Clearance).

· If the HPLC sequence contains a second assay, that assay must be stored in a different .wiff file

The following table is shown for a 5 injection/sample Hepatic Clearance sequence in which samples 1-5 are for Rat and 6-10 are for a second Rat experiment.

	Sample Name
	Data File

	Sample001
	Rat01

	Sample002
	Rat01

	Sample003
	Rat01

	Sample004
	Rat01

	Sample005
	Rat01

	Sample006
	Rat02

	Sample007
	Rat02

	Sample008
	Rat02

	Sample009
	Rat02

	Sample010
	Rat02


Notice that the data file names are different for the two assays.

Note that it is allowable to have several different analytes in each injection (e.g. cassette dosing, looped experiments).

7.2 Thermo

Similar to Sciex data, Thermo data (.raw files) must be grouped by assay. Thermo acquired data cannot be stored in a single data file; therefore, from the perspective of GMSU/QC, the Thermo sequence file (.sld) acts to group Thermo data by assay. The rules vary slightly depending on the CDS used to acquire the data

Note that Thermo applications generate several date-time stamped .sld’s in the course of data acquisition and data processing. GMSU/QC ignores the date-time stamped .sld’s.

· An ‘assay’ is defined as the injections/samples that make up the data set for that analysis (e.g. Hepatic Clearance).

7.2.1 Xcalibur™ data

· An .sld may contain only from a single assay

· If the HPLC sequence contains a second assay, that assay must be configured in a separate .sld and submitted to the sequence queue separately.

The following table is shown for a 5 injection/sample Hepatic Clearance sequence in which samples 1-5 are for Rat. The second table is samples 1-5 are for a second Rat experiment. The two .sld’s would need to be submitted to the acquisition queue separately.

	Sequence Name: Rat01

	Sample Name
	Path

	Sample001
	C:\Xcalibur\Data\Rat01

	Sample002
	C:\Xcalibur\Data\Rat01

	Sample003
	C:\Xcalibur\Data\Rat01

	Sample004
	C:\Xcalibur\Data\Rat01

	Sample005
	C:\Xcalibur\Data\Rat01


	Sequence Name: Rat02

	Sample Name
	Path

	Sample001
	C:\Xcalibur\Data\Rat02

	Sample002
	C:\Xcalibur\Data\Rat02

	Sample003
	C:\Xcalibur\Data\Rat02

	Sample004
	C:\Xcalibur\Data\Rat02

	Sample005
	C:\Xcalibur\Data\Rat02


Please note the following:

· Even though it is permissible to save assay data to the same path, saving assay data to different paths will reduce the possibility of GMSU/QC errors in locating data.

· The two .sld’s must have different names.

· It is allowable to have several different analytes in each injection (e.g. cassette dosing, looped experiments).

7.2.2 QuickQuan™ (QQ) data

QuickQuan is a highly specialized application allowing users to efficiently and quickly setup the instrument for high throughput data acquisition. QQ has an added benefit in that processing methods (.pmd’s) are generated automatically and will be processed automatically by GMSU/QC. See Section 8 for a further discussion of processing methods.

The only limitation with QQ data is that the QQ sequence must be configured such that there is always a 1-to-1 relationship in generated .sld’s-to-.pmd’s.

For example, the following QQ sequence will generate a single data subdirectory with two .sld’s and two corresponding .pmd’s

	New Sequence
	Drug Set
	Sample

Name

	 FORMCHECKBOX 

	QuickCalc
	Sample01

	 FORMCHECKBOX 

	QuickCalc
	Sample02

	 FORMCHECKBOX 

	QuickCalc
	Sample03

	 FORMCHECKBOX 

	Labetalol
	Sample01

	 FORMCHECKBOX 

	Labetalol
	Sample02

	 FORMCHECKBOX 

	Labetalol
	Sample03


In another example, the following QQ sequence will generate two data subdirectories, each containing one .sld and a corresponding .pmd

	New Sequence
	Drug Set
	Sample

Name

	 FORMCHECKBOX 

	QuickCalc
	Sample01

	 FORMCHECKBOX 

	QuickCalc
	Sample02

	 FORMCHECKBOX 

	QuickCalc
	Sample03

	 FORMCHECKBOX 

	Labetalol
	Sample01

	 FORMCHECKBOX 

	Labetalol
	Sample02

	 FORMCHECKBOX 

	Labetalol
	Sample03


7.2.3 LCQuan™ data

GMSU/QC cannot process LCQuan data directly. Therefore, users must first export a desired .sld from within LCQuan. GMSU/QC then may open the exported .sld.

Users must ensure that, within LCQuan, the .pmd is also exported with the .sld.

8 Thermo Processing Methods

There is an advantage to using Thermo data in that GMSU/QC can retrieve information from processing methods (.pmd’s). GMSU/QC pulls as much information as it can from the files available to it from each CDS environment. Table 1 shows a comparison between Thermo and Sciex listing the information that GMSU/QC can retrieve automatically vs information that needs to be configured manually.

Table 1
Thermo/Sciex Automatic Information Retrieval Comparison

	Information

Component
	Thermo 1
	Sciex

	
	Automatic
	Manual
	Automatic
	Manual

	Analyte Name
	X
	
	
	X

	Internal Standard Name
	X
	
	
	X

	Analyte/IntStd Association
	X
	
	
	X

	Compound association with transition ion set
	X
	
	
	X

	Calibration/QC Level
	X
	
	X
	

	Calibration/QC Concentration
	X
	
	
	X


1
With a correctly matching processing method

Though GMSU/QC contains the tools needed to configure manually all the Information Components listed, that action can become quite tedious, especially when throughput is ramped to >50 different analytes/species/assays per day.

If an incorrect (or no).pmd is configured in the Thermo .sld, then all the Information Components must be configured manually.

9 Installation Overview

It is recommended that administrators perform GMSU/QC Setup according to the following steps. Please refer to each referenced section for a detailed discussion of each step

	Installation Step
	Section

	9.1 Inspect system requirements and prerequisites
	Sections 10 - 11

	9.2 Execute Setup.exe
	Section 12

	9.3 Determine the physical location of the GMSU/QC database Gubbs_01.mdb (if applicable)

· Note:  Does not need to be performed if installing a subsequent GMSU/QC installation on an existing GMSU/QC client workstation

· Perform this section if GMSU/QC is to be configured in multi-user environment


	Section 13

	9.4 Prepare Windows® directory share and directory permissions (if needed)

· Note:  Does not need to be performed if installing a subsequent GMSU/QC installation on an existing GMSU/QC client workstation

· Perform this section if GMSU/QC is to be configured in multi-user environment

· Thermo users please pay special attention to Section 15.2

	Sections 15

	9.5 Configure database (if needed)
· Note:  Does not need to be performed if upgrading an existing GMSU/QC installation

9.5.1 Microsoft Access

Does not need to be performed

9.5.2 Microsoft SQL Sever


	Section 16
Section 18

	9.6 Configure DNS ODBC file (if needed)

· Note:  Does not need to be performed if installing a subsequent GMSU/QC installation on an existing GMSU/QC client workstation

9.6.1 Microsoft Access

Perform this section only if GMSU/QC Setup could not automatically configure a DSN ODBC file

9.6.2 Microsoft SQL Sever


	Section 17
Section 19

	9.7 Execute UpdateGMSU/QCDatabase.xls (UGD) (if applicable)

· To be performed with subsequent GMSU/QC installations.
· UGD may or may not need to be executed after a subsequent installation. Upon running the new version, GMSU/QC will notify the user if UGD needs to be executed (see Section 22).


	Section 22

	9.8 Configure GMSU/QC Configuration Default Parameters

· Note:  Does not need to be performed if installing a subsequent GMSU/QC installation on an existing GMSU/QC client workstation

· It is recommended that administrators configure several default values for proper GMSU/QC usage.


	Section 24
Section 26


10 System Requirements

The following describes the environments on which GMSU/QC is configured and developed. The user is free to install GMSU/QC on workstations of differing OS and Microsoft® Office/Excel configuration with the understanding that:

· GMSU/QC should be tested before implementing the differing configuration.

· GMSU/QC may not perform as expected in a lesser configuration than that recommended below.

	Client Workstations

	OS 5
	MS Windows XP SP3

	CPU (GHz)
	1.8

	RAM (MHz)
	512

	Sciex Analyst 1,4
	1.4.1, 1.4.2

	Thermo Excalibur 2,4
	2.0.5

	Microsoft® Office 3, 6
	Office 2003 SP2

	UpdateGMSU/QCDatabase.xls
	MS Excel 2003 SP2

	GubbsIncPswd_01.xls
	MS Excel 2003 SP2

	1
GMSU/QC supports data acquired using Sciex Analyst 1.3.x and Analyst 1.4.2
2
GMSU/QC supports data acquired using XCalibur 2.3 and previous versions

3
At a minimum, Microsoft® Word and Excel must be installed. GMSU/QC is compatible with Office XP, 2003, and 2008 (see note below concerning the PK Calculator 5).

4
Not required for data reporting (see Section 11)

5
GMSU/QC supports and has been tested on the operating system Microsoft Vista, though this operating system has not yet been officially implemented in the development environment.

6
GMSU/QC uses Microsoft Office to generate reports in Excel and Word. GMSU/QC supports and has been tested on Microsoft Office 2008 EXCEPT FOR THE FOLLOWING FEATURE: The PK Calculator has the option of generating a report in Microsoft Word. This feature is not yet supported with Microsoft Office 2008.


	MS SQL Server Server 1

	SQL Server
	SQL Server 2000 SP4, SQL Server 2005

	OS
	MS Windows Server 2003 SP1

	CPU (GHz)
	1.3

	RAM (MHz)
	1250

	1
If SQL Server is to be used


11 Pre-Requisites

11.1 Sciex environment

11.1.1 Data Processing

In order to process Sciex data, Sciex Analyst must be installed on the workstation on which the GMSU/QC client is being installed. GMSU/QC uses several Sciex Analyst Active X objects and .dll’s to ensure that information provided by GMSU/QC is identical to information if it were to be generated on Sciex Analyst.

If Sciex Analyst is not installed on the workstation, GMSU/QC will not allow the user to open Sciex data files (.wiff).
11.1.2 Data Reporting

Sciex Analyst is not required to be installed to report data. All data reporting modules are available.

11.2 Thermo Scientific environment

11.2.1 Data Processing

In order to process Thermo data, the Thermo software development kit called ‘XDK OEM” must be installed on the workstation on which the GMSU/QC client is being installed. GMSU/QC uses several XDK OEM Active X objects and .dll’s.

If XDK OEM is not installed on the workstation, GMSU/QC will not allow the user to open Thermo data files (.sld).

11.2.2 Data Reporting

XDK OEM is not required to be installed to report data. All data reporting modules are available.

12 Setup Execution

NOTES:
If a previous installation of GMSU/QC exists on the workstation, it must be uninstalled (through the MS Windows Add/Remove Programs Control Panel) before a new version of GMSU/QC is installed.

12.1 Setup Execution

Note:  The contents of the screens described in this section may differ slightly from actuality.

Run the GMSU/QC Setup.exe executable. The following screens will be presented:

[image: image1.png]Gubbs Mass Spec Utilities Setup.

g Welcome to the Gubbs Mass Spec Utites instaltion program.

‘Setup cannot nstall system fes or Update shared fies i they are inuse.
Before proceeding, we recommend that you dose any appications you may
berunning.





Click OK

[image: image2.png]2 Gubbs Mass Spec Util
Begin the nstallaton by clcking the button below.

ik this button to install Gubbs Mass Spec Utities software to the specifed
destination drectory.

Drectory:
’Vc Program Fies\Gubbs Mass Spec Utities|





Click the large button with the PC icon. The user may change the location of the installation directory, but it is not recommended.

[image: image3.png]3 Gubbs Mass Spec Utilities - Choose Program Group [X]

‘Setup wil add tems to the group shown in the Program Group box.

You can enter a new group name or select one from the Existng
Groupsist.





This informational window is displayed

[image: image4.png]Version Conflict

A fle being copied is not newer than the fl currently on your
system. It recommended that you keep your existig fle.

File name: ‘C:\WINDOWS system32\comdig32.00x'

Descripton: CMDislog ActveX Control DLL

Your version: '6.1.57.87

Do you want o keep ths fie?

- o




In some instances, the user may be presented with one or more dialog boxes similar to this one. The user should always choose ‘Yes’.

12.2 End User License Agreement (EULA)
The user will be presented with the EULA. If the user does not agree to the EULA, the installation cannot continue.
[image: image5.png]Gubbs Inc GMSU End User License Agreement. (EULA)

End User License Agreement (EULA)

Please readthe following license agreement carsfully.

[GUBBS INC END USER LICENSE AGREEMENT [EULA) FOR Gubbs Mass Spec Utities

BY CLICKING THE "ACCEPT" BUTTON OR OFENING THE PACKAGE, YOU ARE CONSENTING TO BE BOUND BY AND ARE BECOMING & PARTY TO THIS
|AGREEMENT . IF'YDU DO NDT AGREE T0 ALL DF THE TERMS OF THIS AGREEMENT, CLICK THE "D0 NOT ACCEPT"BUTTON. IN THE EVENT YOU DO
INOT ACCEPT ALL THE TERMS OF THIS AGREEMENT, YOU MAY RETURN THE SOFTWARE T0 GUBBS INC FOR A FULL REFUND.

1. THE PARTIES
2] The patis to tislicense are GUBBS INC and the Customer. “GLIBBS INC" means the enfly GUBBS INCORPORATED. The “Custamer” i the person or
erganizaion tha orders and uill pay the licerse for the GMSLI Saftware (‘Saftware,

b} In ranting this inss tothe Customer, GUBBS INC i also acting on behal of Information Pattners. “Information Partners” ae entiies that have censed to
[GUBBS INC programs included s patof the Software. Ech Information Paitner retans s ight o enforce i tademarks, copyights, paterts, iads secrels and
ther ights iectly againstthe Customer or unavtharized capies and urlicensed uses of the Software.

20 LICENSE GRANT
] GUBBS INC grants t the Custome a non-exclusive, norvtansferable ense to use sach authorzed copy of the Software on a single compute system so long as
the Customer complies with the tems o tis Software License Agreemen.

1) The Soltware and the assacialed materas cthet than partians in the public damain) are the propety of GUBBS INC or an Informaion Partner and are protected
by copyright aws,intemational eates. and other ghts

] Notile i tansfered in the Soltware to Customer by ths iense ot by the papment o any fee

) Ay ights ot expressiy granted i his ense e reserved to GUBBS INC and the Information Paitners

] the ficense s tial icense (s displayed upon the fist use of the software),then the tine lmit ofthe fiense i the tim period displayed on th fst use of the.
soltware.

21 WHAT CUSTOMER MAY DO
Customer MaY:

) Use the Software to perfom the functons i was designed for
b) Make one copy ofthe Soltware, for aichival purposes orly.

& {acceptihe terms in the license agreement

I do not aceeptthe terms inthe license agreement





12.3 Query for database type
The user will presented with the dialog boxes in this section only with an original installation on a workstation. If the installation is on a workstation on which GMSU/QC had already been installed, then these dialog boxes will not be presented.

Setup will create a System DSN ODBC file used by GMSU/QC to communicate with the GMSU/QC database. The user will be asked which database type will be used with GMSU/QC.

[image: image6.png]Query for database type...

Please choose the type of database that will be used with this
installation

(Choose Microsoft(R) Access if you don't know)

[ Database type.

@ Microsoft(R) Access

€ MicrosoftR) SQL Server





12.3.1 MS SQL Server

If the user chooses SQL Server, then the following prompt will appear:

[image: image7.png]2

A System DSN ODEC fie called GubbsScex_100' wil b generated.
Please note that your Gubbs Inc GMSU adminsitrator must manually complete the configuration of this DSN fie.





This prompt informs the user that an IT administrator must configure the GubbsSciex_100 DSN file that has been created. See Section 19 for further details.

12.3.2 MS Access

If the user chooses MS Access, the following prompts will appear:

[image: image8.png]0 you wish to configure the DSN with an existng Gubbs_01.mdb that s NOT located at the default Gubbs_01.mdb installation locati of:
C:\GubbsInc|GMSU\Database\Gubbs_01.mdb

-





Choose Yes if there is an existing network-level GMSU/QC database (typically named Gubbs_01.mdb) that exists and it is wished to use that existing database
If Yes is chosen, then the user is prompted to search for the database
[image: image9.png]Browse for shared Gubbs_01.mdb....

If the Gubbs_01.mdb is located on a shared drive, then browse to it now by clicking the Browse... button.

If you don't know the location of the shared Gubbs_01.mdb. then:
1 - Click OK and continue with the default installed Gubbs_01.mdb.
I - Contact your Gubbs Inc GMSU administrator to manually configure the GubbsSciex_100 DSN file.

If you are unable to browse to the location of the shared Gubbs_01.mdb. then:

1 - Click OK and continue with the default installed Gubbs_01.mdb.
I - Contact your Gubbs Inc GMSU administrator to manually configure the GubbsSciex_100 DSN file.

Browse.

CAGubbsinc\GMSLNDatabase\Gubbs_0T mdb,

OK ]| cancel





When the database is found, select it. GMSU/QC Setup will create a DSN file configured to communicate with that database.

[image: image10.png]Select an .mdb file.

5 Database

() Gubbs._01_20070412_101443.md>
() cubbs_01_20070524_171113.md>
() cubbs._01_20070603_160249,md>





If No is chosen at dialog box shown at the beginning of Section12.3.2, then GMSU/QC Setup will create a DSN file configured to communicate with that database installed in the default location (C:\GubbsInc\GMSU/QC\Database\).

NOTE:
If for some reason Setup cannot create the System DSN described in this section, Setup will inform the user that DSN creation was unsuccessful. If this should happen, then the DSN file will need to be created manually according Section 17 for MS Access or Section 19 for MS SQL Server.

12.4 Installation Finish
When the installation is finished, the user will be presented with the following dialog boxes:

[image: image11.png]Visual Basic 6.0

up Tool

Gubbs System DSN Creation Successful




[image: image12.png]up

Gubbs Mass Spec Utites Setup was completed successfuly.

Gubbs Mass Sp





12.5 Additional Actions Performed by Setup
In addition to installing the GMSU/QC software, Setup.exe will perform the following tasks:

12.5.1 Directory creation

Setup.exe will create a C:\GubbsInc\GMSU/QC\ directory and several subdirectories related to GMSU/QC operation. If the directories already exist, then they will not be re-created and existing files within these directories will not be deleted or modified.
The GMSU/QC setup.exe will create the following GMSU/QC directories. 

C:\GubbsInc\GMSU/QC\

C:\GubbsInc\GMSU/QC\Database\ 1,2
C:\GubbsInc\GMSU/QC\Export\ 3
C:\GubbsInc\GMSU/QC\CompoundLists\ 3
C:\GubbsInc\GMSU/QC\ChromatographyMethods\ 4
C:\GubbsInc\GMSU/QC\Manuls\
The folder
…\Database\ is the location in which the GMSU/QC database is originally stored 1,2. GMSU/QC setup.exe places the GMSU/QC database ‘Gubbs_01.mdb’ here.

The folder
…\Export\ is the location in which text file reports will be stored (for Sciex HT Acquisition Tools).

The folder
…\CompoundLists\ is the location in which Automaton™ compound lists are stored (for Sciex HT Acquisition Tools).

The folder
…\ChromatographyMethods\ is the location in which Permeability Calculator chromatographic methods are stored (if chromatographic methods are generated by the user in an HT environment).

The folder
…\Manuals\ is the location in which GMSU/QC Manuals are stored. GMSU/QC setup.exe places current manuals in this directory.

1
If Microsoft Access is to be used as the backend database

2
If MS Access is to be used in a multi-user environment, see Section 15.1
3
Location of temporary .doc and .pdf files when generating PDF reports in the PK Calculator module.

4
Used only by the Permeability Calculator module

Notes:

· If the GMSU/QC database is to be shared by multiple instances of GMSU/QC, it is highly recommended that all local installation GMSU/QC directory structures are not modified once created by GMSU/QC setup.exe. The database stores some local directory structure variables and assumes that all local installation directory structures are identical.

12.5.2 Installation of PDF-XChange® print driver
Setup will install a print driver named ‘PDF-XChange for GubbsInc’. This print driver is used by the PK Calculator module if the user chooses to generate a report in PDF format.

12.5.3 GMSU/QC_01.ini file creation
12.5.3.1 Original Installation
Setup will create a GMSU/QC_01.ini file in C:\GubbsInc\GMSU/QC\.

12.5.3.2 Subsequent Installation

If GMSU/QC_01.ini already exists in C:\GubbsInc\GMSU/QC\. Setup will not overwrite the file.

12.5.4 Gubbs_01.mdb file creation

12.5.4.1 Original Installation

Setup.exe will create a Gubbs_01.mdb file in C:\GubbsInc\GMSU/QC\Database\.

12.5.4.2 Subsequent Installations
If Gubbs_01.mdb already exists in C:\GubbsInc\GMSU/QC\Database\, Setup will not overwrite the file.

13 GMSU/QC Licensing

GMSU/QC is exclusively distributed by Thermo Scientific™. Thermo Scientific manages the licensing of GMSU/QC under the brand Thermo Scientific QuickCalc™.

When a newly-installed GMSU/QC application is installed, the user will be prompted to enter licensing information. Two types of licenses are available for GMSU/QC: Full License and Evaluation License (good for 60 days).

13.1 Licensing Process

The licensing process is as follows:

13.1.1 License prompt

The user will be presented with the following prompt.
[image: image13.png]QuickCale License

Gubbs Inc GMSU (QuickCale) has ot been licensed.

There ate two types of censes avalable:
60 day Evalualion License
Ful Singe License
Flease nole that censes are urique o each workstaton. f you have &
license, please eer it ron.
To obtain aicense, record the fllowing nformation
Your customer nformation (cormpany name, compary st depatment)
This product nformation (DuickCalc)
The type of cense desied Full or Evalation)
The seed license lacated in the License bos below]
and emal his nformalion a ense firigan ems(@thermo,com
‘When the etu icense has been provided, ente it (copy/paste
ecommended) o the License: box below and cick.on the Set button

License: e





Figure 1
License Prompt

13.1.2 Record information
The user must record the following information:

· Customer information:

· Company name

· Company site

· Department

· This product information

· QuickCalc

· The type of license desired

· Full or Evaluation

· The seed license

· Located in the License: text box

13.1.3 Email to Thermo Scientific

The recorded information must be emailed to:


License.finnigan-lcms@thermo.com
13.1.4 Return license key emailed back

Thermo Scientific will respond to the email with a return license key.

13.1.5 Enter the return license

Open GMSU/QC and, at the QuickCalc License prompt, enter the return license.

If the key is not accepted, please:

· Report the problem to License.finnigan-lcms@thermo.com
OR

· Contact your Thermo Scientific representative immediately

14 Physical Location of GMSU/QC database Gubbs_01.mdb

If MS SQL Server is to be used as the GMSU/QC database, then this section may be ignored.

14.1 Stand-alone Environment
If GMSU/QC is to be used as a stand-alone application and the GMSU/QC database will not be shared by other users, then it is recommended that the GMSU/QC database Gubbs_01.mdb be left in its original installation location of C:\GubbsInc\GMSU/QC\Database\.

14.2 Multi-user Environment

In order for GMSU/QC to be used in a multi-user environment, the Gubbs_01.mdb file must be located in a shared directory to which all GMSU/QC users have access. The preferable location for Gubbs_01.mdb is on a network directory, but it can also reside on a local directory that has been shared out to all GMSU/QC users. If a network directory will be used, it is at this point that the network directory should be defined for the next installation step Section 15.

15 Windows Share Permissions and Directory Permissions
The person installing GMSU/QC must be logged onto the workstation with an account that has permissions to modify directory shares and permissions on the local workstation.

If GMSU/QC is to be configured such that multiple instances of GMSU/QC will share the same database and access the same acquired data, a shared directory must be configured according to the following guidelines. This directory can be a directory on a network file server, or the appropriate local directories of a single data acquisition workstation can be shared. Please see Appendix A
Thermo Environment Multiuser System Configuration, for a detailed discussion on configuring GMSU/QC in a multiuser environment.
Please note that Section 15.1 deals mainly with Sciex Analyst users. Thermo data is not typically moved from local data acquisition workstations to network directories (other than for backup procedures).

The following instructions assume that a network file server directory will be configured (Thermo users should see Section 15.2).

15.1 Shared Network Data Directory (if applicable)
Identify a location to which GMSU/QC files will be stored. At this location, generate the following directories:

NOTE:
The easiest way to perform this process is to simply copy a local \GubbsInc\... directory set to the network.

\\[network]\GubbsInc\GMSU/QC\

\\[network]\GubbsInc\GMSU/QC\Database\ 1
\\[network]\GubbsInc\GMSU/QC\Export\ 3
\\[network] GubbsInc\GMSU/QC\CompoundLists\ 2,3
\\[network]\GubbsInc\GMSU/QC\Data\ 2
\\[network]\GubbsInc\GMSU/QC\ChromatographyMethods\ 3,4
The folder
\\…\Database\ is the location in which the GMSU/QC database is stored 1.

The folder
\\…\Export\ is the location in which text file reports are stored.

The folder
\\…\CompoundLists\ is the location in which compound lists are stored.

The folder
\\…\Data\ is the location in which acquired .wiff files are copied. Note that if a network repository of acquired data already exists, that is acceptable. The acquired data directory location is configured as a variable within the GMSU/QC Configuration Utility (for Sciex environment only).
The folder
\\…\ChromatographyMethods\ is the location in which Permeability Calculator chromatographic methods is stored (if a chromatographic method is generated by the user).

1
If Microsoft Access is to be used as the backend database
2
Sciex environment only

3
Optional. If this information will be specific to an instrument, then local directories may be configured in GMSU/QC Configuration (assumes local data directories are configured as described in Section 12.5.1).

4
Used only by the Permeability Calculator module in a HT acquisition environment
15.1.1 Copy files

At a minimum, copy GMSU/QC_01.mdb to \\[network]\GubbsInc\GMSU/QC\Database\.
15.1.2 Share settings

Share the …\GMSU/QC\ directory and configure GMSU/QC users to have a minimum of CHANGE share permissions.

15.1.3 Directory permission settings

Configure directory permissions at the …\GMSU/QC\ and configure the subdirectories to inherit. At a minimum, users must be able to read and write to these directories (MODIFY permissions).

Note: If the directory to which acquired data is stored is something other than a GMSU/QC directory, users must be able to read and write to this directory (Display HT Chromatography acquisition chromatogram text files are written to this directory).

15.1.4 GubbsSciex_100 System DSN ODBC file

Remember that the network Gubbs_01.mdb will be used when configuring the MS Access GubbsSciex_100 System DNS ODBC file described in Section 17.

15.1.5 GMSU/QC clients referencing the shared directory

Use UNC Path

The administrator must set the Configuration Utility – [Module] – Use UNC Path value to ‘Yes’. This will force GMSU/QC to refer and store all references to data files using the UNC path (e.g. \\gubbslap03\Xcalibur\) rather than the drive-letter reference (e.g. C:\Xcalibur\).

15.2 Shared Local Data Directory (if applicable)

GMSU/QC is especially useful when it is configured in a manner that all users have access to the raw chromatographic data (.wiff, .raw, .sld). This type of access is most easily attained when the raw data are stored on a centralized network server. However, it is recognized that the Thermo chromatographic raw data file structure mostly is not amenable to manual movement of data to a network directory. In addition, some Sciex administrators may not wish to include a manual raw data copy step in their data reporting process.
Therefore, if it is desired to use GMSU/QC in a manner that all users have access to local chromatographic raw data, administrators must perform the following steps:

15.2.1 Share the raw data directory (e.g. C:\XCalibur\ or D:\Analyst Data\Projects\)
· The raw data directory must be shared to GMSU/QC users.
· Give the share an intuitive name, such as ‘XCalibur’ or ‘Analyst’

· GMSU/QC users must be granted a minimum of CHANGE share permissions.

15.2.2 Configure raw data directory security permissions
The directory security permissions of the raw data directory must be configured such that GMSU/QC users have a minimum of MODIFY permissions.

15.2.3 GMSU/QC clients referencing the shared directory

Two options are available for all GMSU/QC clients to reference the shared directory

15.2.3.1 Use UNC Path

The administrator may set the Configuration Utility – [Module] – Use UNC Path value to ‘Yes’. This will force GMSU/QC to refer and store all references to data files using the UNC path (e.g. \\gubbslap03\Xcalibur\) rather than the drive-letter reference (e.g. C:\Xcalibur\).

15.2.3.2 Use mapped drive letters

The administrator may map the shared drive on each GMSU/QC client workstation. Administrators must take care that the same letter is used at ALL GMSU/QC clients, INCLUDING data acquisition workstations.
NOTE: Please note in this configuration that, if users are processing data at a data acquisition workstation, they must browse to the MAPPED LETTER when opening data.
15.2.4 Database Location (for MS Access only)

If the data files (.wiff, .sld, .raw) are being stored locally, administrators may still wish to store the GMSU/QC MS Access database Gubbs_01.mdb on a networked file server. If so, Section 15.1 should be followed, except only a directory (e.g. \\[network]\[share]\GubbsInc\GMSU/QC\Database\) need be created and the database copied there.
16 Database Configuration – Microsoft® Access

If MS Access is to be used as the backend database, follow the instructions in this section.

· Setup will place a copy of the Gubbs_01.mdb in the C:\GubbsInc\GMSU/QC\Database\ directory.
· If GMSU/QC Setup could not automatically configure a DSN ODBC file, then configure an appropriate DSN File according to Section 17.
17 DSN Configuration – Microsoft® Access

If Setup.exe could not create the GubbsSciex_100 System DSN ODBC file, then the administrator must manually create the DSN file according to this section.

GMSU/QC communicates with the MS Access database via a system DSN named GubbsSciex_100. Create the DSN according to the following procedure:

17.1 ODBC Data Source Administrator

Open the Data Source Administrator and select the System DSN tab

· Windows XP: via Control Panel – Administrative Tools – Data Sources

[image: image14.png]" ODBC Data Source Administrator.

User SN System DSN | Fi DSN | Drivers | Tracing | Connection Pooling | About

System Data Sources:
Name. Diver

LocalServer SL Server
Xireme Sample Database 2003 Microsoft Access Diver (*mdb]

#in ODBC System data sourcs stores infomation about how to connect to

e e e
it

Cancel Aoy





Figure 2
MS Access DNS - Data Source Administrator

17.2 Create New Data Source

Click on the Add… button to start the Create a New Data Source Wizard, select “Driver do Microsoft Access (*.mdb), and click on Finish

[image: image15.png]e New Data Source

Select  diver for which you want to set up a data source.

Name VA
Drver da Nicrosof para arauivos texo (b4 ~cov) 4

N
‘
‘
N
6
‘
‘
N

Driver do Nicrosof dBase (~dbf)
Drver do Microsoft Excel(xis)
Drver do Microsot Parador (b )
Driver para o Microsoft Visual FaxPro
Microsoft Acoess Drver (mdb)
Microsoft Access-Treber (~mdb)
Microsoft dBase Drver ("db)

<

I





Figure 3
MS Access DNS - Create a New Data Source Wizard

17.3 ODBC Setup

At the ODBC Microsoft Access Setup window, enter the Data Source Name and Description as shown in Figure 4. IMPORTANT: The Data Source Name must be entered exactly as shown.

[image: image16.png]ODBC Microsoft Access Setup

Data Source Name: [GubbsScier_100

Deseiption: [Gubbs Mass Spec Utiiies

Database

Database: CA. \GubbsLlfties\Gubbs_01 s
Select. Ceate.. | Bepai

[

Cancel

Help
Compact

System Database

& Nong

€ Database:

[osntins]|

Advanced

Dptons>>





Figure 4
MS Access DNS - ODBC Microsoft Access Setup

17.4 Select database

Click on the Select… button, navigate to the Gubbs_01.mdb, and click on OK

[image: image17.png]Select Database.

Database Name [E—

[oos i clanayet dete

[=C

5 Anaiyst Data
B GubbsLilties
23 CompoundLists
0 PlateMaps

Lt Fles of Type:
[Access Databases ('m ~





Figure 5
MS Access DNS - Select Database

Click on the OK button to complete the DSN configuration process.

18 Database Configuration – Microsoft® SQL Server

If Microsoft SQL Server is to be used as the backend database, follow the instructions in this section.

18.1 Create a SQL Server database using the MS Access Upsizing Wizard
Setup will place a copy of the Gubbs_01.mdb in C:\GubbsInc\GMSU/QC\Database\.

Open this file in MS Access and run the Upsizing Wizard

[image: image18.png]e Edt Vew Insert

HRR= A= RESFE NS





Figure 6
MS Access Upsizing Wizard
The following figures show the parameters to choose when performing the Upsize action

[image: image19.png]The Upsizing Wzard alows you to easly upsize your Microsaft Access
database to 3 Mirosoft 5L Server database.

You can create 3 new SQL Server database or upsize to an existing
5QL Server database, Which do you wart ta do?

€ Use existing detabase

& Greateow daisbess |

after you cick the Next button, youll supply nformation for the new
database.

Help Concel ek | wexts Eish





Figure 7
MS Access Upsizing Wizard - Choose Create new database

[image: image20.png]What SQL Server would you ke to uss for this database?

EE 2

Please specfy the ogin 1D and passward of an accaunt
with CREATE DATABASE privieges on this server.

T~ Use Trusted Connection

LognDD;  [qubbs

passuord: [

What do you want to name your new 5QL Server database?

[oubbs_o1saL

Help Concel <gack | wext> Eish





Figure 8
MS Access Upsizing Wizard - Connection settings

Here, enter a login ID and password of an account with sufficient permissions to create a database (e.g. sa). A GMSU/QC user account will be generated at a later step. The database name can be of your choosing.

[image: image21.png]up

g Wizard

Which tables do you wart ta exgort to QL Server?

Avalable Tables: Export to SOL Server

tblatchTemp
5 | [icae
thiCacochromPeakiiames
> | |ticacaconstants
thiCaccRecsum
< | |tblcacosummry
thiCacosummaryamples
<< | [biCalircaco
thiChvomirethodiiame
thichvomiethodvalue

Selected Tales [EPutomatonChange

= =] e e | Lae





Figure 9
MS Access Upsizing Wizard - Export all tables

[image: image22.png]up

g Wizard

The Upsizing Wizard can export table atributes n additon to data.

Wht table attributes do you wan to psize?
W findeves ¥ pefauks
W Valdation rues ¥ Table relationstips

& UseDRL € Usetriggers

What data optons do you want to incude?

‘Add timestarp fields to tables?; [N, never 2

T only create the table structure; don't upsize any data.

= =] e e | Lae





Figure 10
MS Access Upsizing Wizard – Export attributes

[image: image23.png]up

g Wizard

The Upsizing Wizard can modfy the existing applcation o create a new applcation ta work
with the SQL Server database.

Wht pplcation changes do you wart to make?

€ reate a new Access clentsever applcation.

£DPFiefieme: [EGUBbSInGMS GBS OICs.adp | _Bpovse
 LinkSQL Server tables to existing application

o appiication changes.

I Saye passyiord and user 10;

Help Concel

<gack | wext> Erish





Figure 11
MS Access Upsizing Wizard – No application changes

[image: image24.png]The Upsizing Wrzard has al the nformation : needs to upsize
your database.

Help Concel <gack | e Erish





Figure 12
MS Access Upsizing Wizard - Finish

At this point the Upsizing Wizard will generate a report which may be discarded.

18.2 Create a GMSU/QC database account

In SQL Server, create a GMSU/QC database account. The account name and password can be of your choosing.

Add the account to the GMSU/QC database and assign the account the following ROLES:
[image: image25.png]Gensal |

s G

User nare: Gubbs

Database role membership:

[Permit in Database Fole =

db_owrer
db_accessadnin
db_secuityadinin
db_cdadmin
db_backupsperator
do_datareader
dbo_ctawiter
db_deydatereader
db_derydatawiter

Dooooooo

|
Propeties

oK Cancel foply || Heb





Figure 13
MS SQL Server GMSU/QC account roles
18.3 Record the user account and password information
Record the user account and password information on a sheet of paper. The user account and password information is needed to update the GMSU/QC_01.ini file described in Section 20.
18.4 Local Data Source Name (DSN)

GMSU/QC communicates with the SQL Server database via a system DSN. Create the DSN according to the procedure described in Section 19.
19 DSN Configuration – Microsoft® SQL Server
If MS SQL Server is to be used as the backend database, follow the instructions in this section.

If Setup.exe could not create the GubbsSciex_100 System DSN ODBC file, then the administrator must manually create the DSN file according to this entire section.

If Setup.exe was successful in creating the GubbsSciex_100 System DNS ODBC file, then the administrator must complete the configuration starting with Step 19.3 of this section.

GMSU/QC communicates with the SQL Server database via a system DSN named GubbsSciex_100. Create the DSN according to the following procedure:

19.1 ODBC Data Source Administrator

Open the Data Source Administrator and select the System DSN tab

· Windows XP: via Control Panel – Administrative Tools – Data Sources

[image: image26.png]" ODBC Data Source Administrator.

User SN System DSN | Fi DSN | Drivers | Tracing | Connection Pooling | About

System Data Sources:
Name. Diver

LocalServer SL Server
Xireme Sample Database 2003 Microsoft Access Diver (*mdb]

#in ODBC System data sourcs stores infomation about how to connect to

e e e
it

Cancel Aoy





Figure 14
SQL Server DNS - Data Source Administrator

19.2 Create new data source

Click on the Add… button to start the Create a New Data Source Wizard, select “SQL Server”, and click on Finish

[image: image27.png]Create New Data Source.

Select  diver for which you want to set up a data source.

Name VA
Microsoft Paradox-Treber {.db )

Microsoft Text Drver (“bé; “csv)

Microsoft TextTreber (b4; csv)

Microsoft Visual FoxPro Driver

Microsoft Visual FoxPro-Treer

Oracle in OraDb 10g_home

SQL Native Client

QL Server

Themo Bleciron DataDirect Oracle Driver

<





Figure 15
SQL Server DNS - Create a New Data Source Wizard

19.3 Create a new data source
At the Create a new data source window, enter ‘GubbsSciex_100’ as Name and Description as shown. Choose the server on which the GMSU/QC SQL Server database was generated, then click Next

IMPORTANT: The Data Source Name must be entered exactly as shown.

[image: image28.png]Create a New Data Source to SOL Server.

“This wizard wil help you creste an ODBC data source that you can use to
comnectto SQL Server.

What name doyou wart fo use o referto the data source?

Neme: [GubbsSciex_100 — |
How do you want o descrbe the data source?

Descipton: [GMSU SGL Server Connection

Which SQL Server do you want to connect t0?

I





Figure 16
SQL Server DNS - Create a new data source
19.4 Enter connection parameters
Configure the next window as shown. Enter the userid and password combination created in Section 18.2. Then click Next.
[image: image29.png]Create a New Data Source to SOL Server.
How should SQL Server vrfy the authentity of the login ID?

€ Wt Windows NT authenticatio using the network logn ID.

s Wih SOL Server athenticaton using  logn 1D and password
entered by the user

To change the network brary used to commuricate wih SOL Server,
cick Clert Corfiguraton.

et Configuration.

Connectto SQL Server to obtain defauk settings for the.
‘addtional corfiguraton options.

LogniD: fokbs
e

s [T ] _om |





Figure 17
SQL Server DNS - Enter connection parameters
19.5 Additional parameters

Enter additional parameters as shown in the next two figures, then click Next and Finish:

NOTE: Ensure that the default database is the GMSU/QC database.
[image: image30.png]Microsoft SQL Server DS Configuration

7 Biange the geaul database 1]
[Gubos_015QL_new
I~ Atach detabase iename.

pttpetess

eate temporay stored procedires forprepered SAL staerents
W7 e chop the stored procedures:

& Drlp when you discannect

when you discorrect and =5 appropiats whie you at
Gonnected.

¥ Use ANS quoted identfiers.
¥ Use ANSI nuls, paddings and warings.

1 Usetfefofover S Syt pinay SO Sy ot

O o





Figure 18
SQL Server DNS - Enter additional parameters

[image: image31.png]Create a New Data Source to SOL Server.

Use strong encryption for data

Berfomtranslation for character data
Use regional setings when outputting curency, numbers, dates and
fimes.

‘Save long running queres to the log fie:
T Biowse,

Long query tine (riliseconds):
Log QDBC diver statisics o thelog e





Figure 19
SQL Server DNS - Enter additional parameters

19.6 Test connection

Test the connection by clicking on the ‘Test Data Source’ button
[image: image32.png]ODBC Microsoft SOL Server Setup

Anew ODBC data source wil be created with the folowing
‘configuraton

[Microsoft SGL Server ODBC Diver Version 03.85.1117

|Data Source Name: GubbsSciex_100

|Data Source Descrpton: GMSU SQL Server Connection
|Server: GUBBSD4

[Database: (Defaut)

|Language: (Defaut)

[Translate Character Data: Yes

|Log Long Running Queres: No

|Log Drver Statitcs: No

Use Integrated Securty: No

Use Regonal Setings: No

[Frepared Statements ption: Drop temporary procedures on
ldisconnect

Use Falover Server: No.

Use ANS| Quoted Idertfiers: Yes

Use ANSI Nul, Paddings and Wamings: Yes

[Data Enciyption: No

ox_| _coea





Figure 20
SQL Server DNS - Test connection
[image: image33.png]Test Resuls

[Microsoft SQL Server ODBC Drver Version 03.85.1117
[Runring connectiviy tests.

|Atempting connection

[Connection estabished

|Verfying option seftings

| Disconnecting from server

ITESTS COMPLETED SUCCESSFULLY!





Figure 21
SQL Server DNS - Test connection

If the last line does note read “TESTS COMPLETED SUCCESSFULLY”, then contact your GMSU/QC administrator before proceeding.

Click on the OK button to complete the DSN configuration process.

20 GMSU/QC_xx.ini

Setup will create an .ini file (named GMSU/QC_xx where ‘xx’ is a number) located in the directory C:\GubbsInc\GMSU/QC\.

Please note that if the original client GMSU/QC installation version was prior to GMSU/QC 2.0.50, then the DSN line in the client .ini file will not exist. If it is desired to configure test instances of GMSU/QC according to Section 23, then the DSN line must be entered manually to the .ini file according to the following procedure:

· Open GMSU/QC_01.ini in Notepad or a suitable text file reader (MS Word is NOT recommended)

· Add a line after the Pswd line

· Enter DSN

· Enter a tab

· Enter GubbsSciex_100

· Save then close the file

20.1 MS Access GMSU/QC_xx.ini configuration

The default .ini content is for a MS Access installation. Please note that the UserID and Pswd entries of the .ini file may be ignored. Ensure that the .ini file has the following content:

boolAccess
True

boolSQLServer
False

UserID
[USERID]

Pswd
[PASSWORD]

DSN
GubbsSciex_100

Note:
UserID is ignored


Pswd is ignored

20.2 SQL Server GMSU/QC_xx.ini configuration

If SQL Server is to be used as database, ensure that the .ini file has the following content:

boolAccess
False

boolSQLServer
True

UserID
[USERID]

Pswd
[PASSWORD]

DSN
GubbsSciex_100

Note:
[USERID] is the SQL Server account generated in Section 18.2.


[PASSWORD] is the password of the SQL Server account generated in Section 18.2
IMPORTANT: The Gubbs Inc Password Encryption Tool described in Section 21 MUST be used to encrypt the password. Copy the encrypted password form the Password Encryption Tool into the Pswd portion of the .ini file (remember to include the trailing space in the encrypted password obtained from the Encryption Tool).
21 Gubbs Inc Password Encryption Tool

Setup will place a file named GubbsIncPswd_01.xls in the installation directory. Open this file to display the Gubbs Inc Password Encryption Tool. Copy the encrypted password to the .ini file if necessary (remember to include the trailing space in the encrypted password obtained from the Encryption Tool).
[image: image34.png]Gubbs Inc Password Encryption Tool 1.0

Enter Password:

Encrypted Password:

Copy entire es
Exit

ncrypted password, including trailing space





22 UpdateGMSU/QCDatabase.xls (UGD)
UpdateGMSU/QCDatabase should be run when a subsequent version of GMSU/QC is installed on an existing GMSU/QC client workstation. If GMSU/QC database is part of a multi-user environment, UGD needs to be executed only once. UGD does not need to be run when the system is initially installed and configured.

If it is unknown if UGD has been executed, simply open GMSU/QC. GMSU/QC will prompt the user if UGD needs to be executed.

[image: image35.png]3 GMsU Console

Gubbs* Mass Spec Utilities | About..|
_tep...|

Dude. UpdateGMSUDatabase xIs (UGD) must be executed before GMSU can be opened.

You may start this file through the
Windows Start - All Programs - Gubbs Mass Spec Utilities menu

Exclusively Disibuted By ‘Watkstation Prerequisies Installedt A product o Gubbs Inc

[V Thema Scientific(R) XDK-OEM
Thermo S s $LVERT

SCIENTIFIC





Figure 22
GMSU/QC UGD notification

NOTE: Ensure that the workstation Excel Macro Security (Tools – Macro – Security) is set to Medium. If Macro Security is set to High, then UGC will not run.

Setup will create UpdateGMSU/QCDatabase.xls in the installation directory. UGD may be accessed through the Windows Start – All Programs – Gubbs Mass Spec Utilities menu:
[image: image36.png]Gubbs Mass Spec Utites
PDF-¥Change 3
BOF-XChance PRO 3 SDK.





22.1 Description of UGD contents

When UGD is opened, the user will see two worksheets. The first worksheet named ‘Sheet 1’ contains the Execute Update button used to execute the database update (see Figure 23).
[image: image37.png]B3 Microsoft Excel - UpdateGMSUDatabase.xis

) Fle Edt Vew Insert Fomat Iods Data Window Hep AdobePOF Type a question for help
DERRASRTHREISDB I ] FO R T |

L2820 0D 0) 5 o) 121 B A YRy it cranges.. exdreven. I

imm v =i =B Z U g Biv o scamy. |

1
[ the database backend is Microsoft Access, then the database Gubbs_01.mdb will be backed up automatically before the actual Update actions occur.
The backup will be given a name appended with a datetime stamp. E.G.. ...\Gubbs_01_20070412_100955.mdb
I the database backend is Microsoft SQL Server, itis STRONGLY advised that a system administrator backup the database before executing this Update

Execute Update

S

o
3

e |60 LG G |

o ¢ > *\Sheet1 ini Path
Draw~ [3 | Autoshapes- N\ N 1O ~| dl < (8] ] &~ F - A
Ready





Figure 23
UGD Worksheet ‘Sheet 1’
The second worksheet named ‘ini_Path’ contains the path of the .ini file that contains the name of the DSN file used to connect to the database (see Figure 24). See Section 20 for a description of the GMSU/QC .ini file.

[image: image38.png]B Microsoft Excel - UpdateGMSUDatabase.xls.
] He Edt Vew Insert Fomat Toos Data Window tep AdobePDF

DEHRSGRIZEE % LB 19018 813 wBw -ef

LR D) 5 o) 213 A YRy it cranges.. Exdreven. I

Em v ERERE = W HS % o

Type a question for help

Biv o scamy. |

- 2

R

7

8

9

= Browse.

C:\Gubbsinc\GMSU\GMSU_01.ini

=
=
9
5=}
-
&
E
4
El

N Sheett ) ini_Path
Duaw~ s | Autoshapes- N\ N 1O Al &1 8 & &2 A
Ready





Figure 24
UGD Worksheet ‘ini_Path’

22.2 Procedure

22.2.1 Open UGD

22.2.2 Select the sheet ‘ini_Path’ and ensure ini_Path is correct
The default entry for ini_Path is the path and file created by Setup. The administrator must modify this entry if any of the following is true:

· The administrator has moved the .ini file to a different location

· The administrator has configured several instances of GMSU/QC databases (e.g. for testing purposes) according to Section 23.

22.2.3 Select the sheet ‘Sheet 1’ and click on the Execute Update button

After clicking the Execute Update button, the date of execution and Notes about the update event will be record in ‘Sheet 1’

[image: image39.png]K3 Microsoft Excel - UpdateGMSUDatabase.xls.
] He Edt Vew Insert Fomat Toos Data Window tep AdobePDF
DEEHRISRTEE 8 DB F)9 O
L2263 y
1 =B Z U !..sﬂ.m,...\gg»m

Type a question for help

[ the database backend is Microsoft Access, then the database Gubbs_01.mdb will be backed up automatically before the actual Update actions occur.
The backup will be given a name appended with a datetime stamp. E.G.. ...\Gubbs_01_20070412_100955.mdb

If the database backend is Microsoft SQL Server it is STRONGLY advised that a sfﬂem administrator backuf the database before execu\m? this UEdale

Execute Update

SITWET

Date Executed: June 21, 2007 08:02:34

Notes: Database backed up successfully as:C-\Analyst Data\GubbsUtiities\Gubbs_01_20070621_080234.mdb
Field numAnalyteP! already exists in tbIPKCalculator . No need to update.
Field numAnalyteDI already exists in tbIPKCalculator . No need to update.
Field numISP! already exists in tbIPKCalculator . No need to update.
Field numISDI already exists in thIPKCalculator . No need to update.
Field numTotalTrans already exists in tbIPKCalculator . No need to update.
Field numTransSetNum already exists in tolPKCalculator . No need to update.

e LG ELIER G|

Draw~ g | Autoshepes= N\ N (1O Al &1 I8
Ready





Figure 25
UGD – After Execute Update
If a database update event has occurred, then the SQL statement used to generate the change will be shown

[image: image40.png]3 Microsoft Excel - UpdateGMSUDatabase.xls

] He Edt Vew Insert Fomat Toos Data Window tep AdobePDF
DEEHRISRTEE 8 DB F)9 O
|36 ¥

-1 s B Z U

T Biv o scamy. |

If the database backend is Microsoft Access, then the database Gubbs_01.mdb will be backed up automatically before the actual Update actions occur.

The backup will be given a name appended with a datetime stamp. E.G.. ...\Gubbs_01_20070412_100955.mdb

If the database backend is Microsoft SQL Server it is STRONGLY advised that a sfﬂem administrator backuf the database before execu\m? this UEdale

Execute Update

SITWET

Date Executed: June 21, 2007 08:02:34
“alter table tbiPKCalculator ADD numSampleOpt short NULL" successfully executed on thIPKCalculator.
“alter table thiPKCalculator ADD charCDS text(100)" successfully executed on thiPKCalculator.
“UPDATE thIPKCalculator SET numSampleOpt = 3 WHERE numSampleOpt IS NULL" successfully executed on tblPKCalculator.
“UPDATE thlPKCalculator SET charCDS ='ABI' WHERE charCDS IS NULL" successfully executed on tbIPKCalculator.
“CREATE TABLE tblGenericViewer (id_tbiGenericViewer int CONSTRAINT PK_id_tblGenericViewer PRIMARY KEY)," Table thiGenericViewer cre:
“alter table thiGenericViewer ADD charAnalyte text(100)" successfully executed on tblGeneric\Viewer.
“alter table tbiGenericViewer ADD charblethod text(100)" successfully executed on tblGeneric Viewer.
uccessfully executed on thiGeneric Viewer.

e LG ELIER G|

Draw~ g | Autoshepes= N\ N (1O Al &1 I8
Ready





Figure 26
UGD – SQL Statements

22.2.4 Post execution review

If UGD experienced any problems during the database update, then that event will be shown in red font and the admin will be prompted to contact Gubbs Inc.

The administrator should review the results of Sheet 1 and look for red font and report any anomalies to Gubbs Inc

If the administrator wishes to save the results of the update, he/she should save UGD using SAVE AS and rename the Excel file.
23 Configuring Test and Production GMSU/QC Database Instances

Administrators have the ability to configure different instances of GMSU/QC databases. This would be useful if users wish to maintain, for example, a testing database and a production database.

23.1 Procedure

The followings steps describe the procedure for configuring a testing database.
23.1.1 Copy GMSU/QC database table.

Create a copy of the GMSU/QC database table and give it a descriptive name (e.g. GMSU/QC_TEST). If MS Access, it is suggested to place the copy in C:\GubbsInc\GMSU/QC\Database\.

23.1.2 Create a new GMSU/QC DSN file

Create a new GMSU/QC DSN file according to Section 17 (for MS Access) or Section 19 (for MS SQL Server), except do the following:

· Point the new DSN file to the new database copy (e.g. GMSU/QC_TEST)

· Name the new DSN file something different (e.g. GMSU/QC_TEST)

23.1.3 Copy GMSU/QC_01.ini

· Create a copy of the GMSU/QC_01.ini file (it suggested to leave it in C:\GubbsInc\GMSU/QC\)
· Name the file something descriptive (e.g. GMSU/QC_TEST.ini)

23.1.4 Modify the new .ini file

· Open the new .ini file in Notepad

· If the .ini file does not contain a last line of:

DSN
GubbsSciex_100

then create the line (Note: DSN and GubbsSciex_100 are separated by a tab)

· Change the DSN entry to the DSN name created in Section 23.1.2. E.g.

DSN
GMSU/QC_TEST

(Note: DSN and GMSU/QC_TEST are separated by a tab)

23.1.5 Create a shortcut

On the Windows desktop, right-click and choose New - Shortcut
[image: image41.png]ATI CATALYST(R) Contral Center

Arrange lcons 8y
Refresh

paste
Paste Shorteut
Undo Rename.

sz

(3 simep tnage
Datasharks Donrloader Session

8] Microsoft Word Document

3] Microsoft Offce Access Appication

) Microsoft roject Doaument

] MirosoftPonerpoit Presentation
T Adobe Photoshop Inage

{28) Microsoft Offce Publsher Document
WY et Secuence

A sqL server Profiler - trace data fle

) microsoftiso Drawing

@ wav Audo

) microsotExce orisheet
(3 Compressed (Zpped) Folder





23.1.6 Browse to GubbsMassSpec.exe

At the browse window, browse to C:\Program Files\Gubbs Mass Spec Utilities\GMSU/QCtilities.exe

[image: image42.png]This wizard helps you to create shortauts tolocal or
network programs, fies, folders, computers, or Internet
addresses.

Type the location of the item:

Clck Next to continue.





[image: image43.png]Browse For, Folder

Select the target of the shortcut below:

= 5 Gubs Mass Spec Usites
@ 7o
@ ) Conpoundust 00016t
@ [ usu_otin
® ) Gusu_Admnanial 0002.doc

& T GSU_Adminvanus_0002 5t
& ] Gsu_userMansl05.doc
% T GSu_UserManusl 05,pcf





Click OK

[image: image44.png]This wizard helps you to create shortauts tolocal or
network programs, fies, folders, computers, or Internet
addresses.

Type the location of the item:

o s Spec Unioes Gt e

Clck Next to continue.





Click Next

23.1.7 Name the shortcut

Name the shortcut (e.g. GMSU/QC_TEST)

[image: image45.png]Select a Title for the Program.

Type aname for this shortaut:
GSU_TEST

Clck Fnish to create the shortaut.





Click Finish. A shortcut named GMSU/QC_TEST will appear on the desktop.
[image: image46.png](GMSU_TEST




23.1.8 Modify the shortcut target
· Right-click on the shortcut and choose Properties
· Select the Target field and move the cursor to the end of the text

· Type a space and a quotation mark

· Type the path and name of GMSU/QC_TEST.ini

· Type a quotation mark

· Click OK

[image: image47.png]GMSU_TEST Properties

General | Shortcut | Compatbity | Securty |

& GMs_TEST

Target type:
Target location:

"C:\Program Fles'\Gubbs Mass Spec Utiies”

[None





By clicking on the shortcut, GMSU/QC will load the GMSU/QC_TEST database instance.
24 GMSU/QC Configuration Default Parameters

Please note that this section may be ignored if the GMSU/QC installation is a subsequent installation on an existing GMSU/QC client workstation.

Several GMSU/QC configuration parameters must be configured in the Configuration Utility in order that GMSU/QC functions properly upon initial use. These parameters are defined in the following sections of the Administration Manual.
Note:
If the Configuration Utility Section item in the table below will not be used, then the configuration section may be skipped.

	Configuration Utility Section
	Administration Manual Reference

	Caco (for Permeability Calculator Module)
	Section 26.1

	Generic Viewer (for Generic Chromatographic Viewer Module)
	Section 26.6

	Global
	Section 26.7

	Hepatic Clearance
	Section 26.8

	PK Calculator
	Section 26.13


25 Console
The GMSU/QC Console is the module from which users navigate to other modules. The Console contents may differ slightly from what is shown.
The Workstation Prerequisites Installed frame box will show the system prerequisite components currently installed on the system. If only Thermo components are found, then the Sciex HT Acquisition Tools item will not be shown.
[image: image48.png]3 GMsU Console

Gubbs* Mass Spec Utilities

- Home Sciex® HT Acquisition Tools

Hepatic Clearance Calculator
Permeability Calculator

Generic Chromatographic Viewer
PK Calculator

Configuration Utility

Exclusively Disrbuted By ‘Watkstation Prerequisies Installedt

Thermo ¥ Themo Scientfc(F) XDK-QEM

ST Ne % Scies(R) Anast ]

A product o Gubbs Inc

SLVERT





Figure 27
GMSU/QC Console

Users click on the appropriate buttons either to open a utility or to navigate to an additional hierarchical level.

26 Configuration Utility
It is recommended that administrators and/or power users peruse all the options listed in the Configuration Utility. The utility sections contain many parameters that are configurable by the user.

Administrators use the Configuration Utility to configure such things as global variables, contents of module dropdown lists, and units (e.g. mg/mL) of different parameters. Users choose the appropriate module to configure by choosing an item in the ‘Choose a Category’ list. 
The initial view of the Configuration Utility is a non-editable mode. In order to edit a value, the user must select the value, then click the Edit button. The user will be prompted to enter a new value, after which the user will click the Save button.

26.1 Caco - Global

The Caco – Global window contains default paths and values specific to the Permeability Calculator. The parameters must be configured in order for proper GMSU/QC operation follow the figure
[image: image49.png]#2 Configuration Utility.

Choose a category:
[
Flags

Generic Viewsr

Giobal

Hepalic Clearance

Hepalic Clearance Constan
Moy Method

PK Calculator

‘Anslysis Time Points

Chromatography Peak Sets

Constarts

Description

Default Bunching Factor Value

Default Caco Constants Set

Default Chromatography Method

Default Chromatography Method Fath

Default Data Directory

c

Default Data File Access Mode

2 - Open all data files within a directory

Default Peak Name Set

Default

Default Results Plot Regression Significant Figures

s

Default Results Significant Figures

B

Default Smoothing Factor

1

Default Time Point Set

Excel template file

LT or HT Data Acquisition?

Shovs Low Throughput/High Throughput Option Windowy

Use IS as Default Setting





Figure 28
Configuration Utility – Caco (Permeability Calculator) – Global

26.1.1 Initial use parameter configuration
The following are parameters that must be configured upon initial GMSU/QC installation.

Note:
If this module will not be used, then this section may be ignored.

26.1.1.1 Default Method Chromatography Path

· This directory is where Permeability Chromatographic Method text files (for HT Acquisition) are stored.
· This directory is created by GMSU/QC Setup (see Section 12.5.1):

C:\GubbsInc\GMSU/QC\ChromatographyMethods\
26.1.1.2 Default Data Directory

· This directory is the root directory of where raw data is stored. This directory is shown initially when the Browse… button is clicked in the Permeability Calculator module.

26.1.1.3 Show Low Throughput/High Throughput Option Window

· Yes/No to view the LT/HT option window in the Permeability Calculator module.

· If both LT and HT acquisition data (see Section 3 for a description of LT and HT acquisition) will be viewed by users, then set this value to Yes.

· If not, set this value to No

26.1.1.4 LT or HT Data Acquisition

· Set value to LT if only LT acquisition data will be viewed.

· Set value to HT if only HT acquisition data will be viewed.

· This value is ignored if ‘Show Low Throughput/High Throughput Option Window’ is set to Yes

26.2 Caco – Analysis Time Points

The Analysis Time Points window is used to create Time Point Sets that will be assigned to analysis data files. ‘Time Points’ are defined as the assay time points in minutes.
Users click on the Edit button to configure a new Time Point Set. Users click on the Save button to actually create the Time Point Set. Once a Time Point Set has been created, it cannot be modified.

[image: image50.png]#2 Configuration Utility.

E= B = S

Choose a category:

Flags
Generic Viewsr
Giobal

Hepalic Clearance

Moy Method
PK Calculator

Constarts
Hepalic Clearance Constan

Global

‘Analysis Time Points

Chromatography Peak Sets

View a Time Point Set
(306090120
Set_204060





Figure 29
Configuration Utility – Caco (Permeability Calculator) – Analysis Time Points

26.3 Caco – Chromatography Peak Sets
The Chromatography Peak Sets window is used to create Chromatography Peak Sets that will be assigned to analysis data files. These sets provide identities of the chromatographic peaks of the acquired data.
Users click on the Edit button to edit an existing Chromatography Peak Set or configure a new Chromatography Peak Set.

Peak identities are assigned by selecting the appropriate Peak Name in the Peak Name list, then double-clicking the appropriate row in the 

Users click on the Create New Set button to configure a new Chromatography Peak Set. Users click on the Execute button to actually create the Chromatography Peak Set, then click the Save button to save the set.

[image: image51.png]#2 Configuration Utility.

‘Anslysis Time Points

Chromatography Peak Sets

Enter total number of chromatographic peaks:

Enter number of calibration levels:
Choose an Analysis Time Point Set:

Enter a Comment (Optional):

Double-clickto set Peak Name.
Chromatography Peak Set

Constarts

B
2

A5t T
AS1d2
A543
ASide
ASd5
DA

DA

DA

DAGD
DAGD
DAGD

Blerk
RE15

RE1S

RE1S
REA0
RE30
REA0
RE45

RE45





Figure 30
Configuration Utility – Caco (Permeability Calculator) – Chromatography Peak Sets (shown in Edit mode)
26.4 Caco – Constants

The Constants window is used to create Constants Sets that will be assigned to analysis data files. These sets include the various constants used in the calculation of Permeability Experiments.

Users click on the Edit button then Create New Set button to configure a new Constants Set. Users click on the Save button to actually create the Constants Set. Once a Constants Set has been created, it cannot be modified.

[image: image52.png]#2 Configuration Utility.

E= B = S

Choose a category:
e |
Flags nalss Tine Poits Chromstography Peak Sets
(Gerei Viewsr
Global
Hepatc Clearance
Hepatc Clearance Constan Ay
Modiy Method cekalbecrg i)
PK Calcualor Basal Dosig nl)
Concenliaton uMole/L]
Growth Area 2]
Semping fom A (nL)
Samping om ()
Dosing Soluion Volure ()
Dition Vokume (rL)

Constants





Figure 31
Configuration Utility – Caco (Permeability Calculator) – Constants

26.5 Flags
The Flags window configures the contents of the ‘Chromatogr. Review’ dropdown box of the Hepatic Clearance Calculator. ‘Results Flags’ are not configurable.

[image: image53.png]#2 Configuration Utility.

E= B = S

Flag Chromatography Flags
(Geneic Viewsr
Global No Flag

Hepalic Clearance Cool Results
Hepalic Clearance Constan
e Erratic results

P Coluictor Low sensitivity

Poor chromatography
Sample Prep Problems
Suspicious solubility

Results Flags
ot Reviewed
[Acceptable
Goofy
Unacceptable





Figure 32
Configuration Utility - Flags
26.6 Generic Viewer

[image: image54.png]#2 Configuration Utility.

Choose a category:
Caco
Fiags

Giobal
Hepalic Clearance

Hepalic Clearance Constan
Moy Method

PK Calculator

Description

Value

Default Bunching Factor Value

1

Default Column Heading Template

Default

Default Data Directory

c:\

Default Data File Access Mode

1- Open a single data file within a directory

Default Results Decimal Places

2

Default Results Plot Regression Significant Figures

s

Default Results Significant Figures

B

Default Results Use Sig Figs (Yes) or Decimals (No)

Default Smoothing Factor

Excel template file

Flag Criteria (%) - Analyte Retention Time

10

Flag Criteria (%) - Int Std Peak Area

20

Flag Criteria (%) - Int Std Retention Time

10

Flag Criteria (%) - QC Peak Area

20

Sample Number to Evaluate Chromatography

B

Select Default CDS Environment

Thermo(R)

Show biinking flag labels

Yes

Use 15 as Default Setting

Yes

Use UNC Path

Yes





Figure 33
Configuration Utility – Generic Viewer

The Generic Viewer window contains global parameters concerning default directory paths of system components.

26.6.1 Initial use parameter configuration

The following are parameters that must be configured upon initial GMSU/QC installation.

Note:
If this module will not be used, then this section may be ignored.

26.6.1.1 Default Data Directory

· This directory is the root directory of where raw data is stored. This directory is shown initially when the Browse… button is clicked in the Generic Chromatographic Viewer  module.

26.6.1.2 Select Default CDS Environment

Used only if users will be viewing both Thermo and Sciex data. This setting governs the order in which ‘File Type Filter’ are displayed when the Browse… button is clicked.

· If the setting is ‘Thermo(R)’, then the default choice is ‘.sld’.

· If the setting is ‘ABI(R)’, then the default choice is ‘.wiff’.

26.6.1.3 Use UNC Path

If data acquisition files (.wiff or .sld) are to be stored at the local data acquisition workstations AND the administrator wishes to configure GMSU/QC in a multi-user environment, then Use UNC Path must be set to ‘Yes’. Please see Appendix A
Thermo Environment Multiuser System Configuration, for a detailed discussion on configuring GMSU/QC in a multiuser environment.
26.7 Global

The Global window contains global variables concerning default directory paths of system components.
[image: image55.png]#2 Configuration Utility.

Lt | e | o] et |

Hepaic Clearance
Hepaic Clearance Consten e
Modiy Method Descgtion]
P Coluictor [Analyst Acquisition Data Path
Assay Request File Path

Report Text File Export Location





Figure 34
Configuration Utility - Global

The Global window contains global parameters concerning default directory paths of system components
26.7.1 Initial use parameter configuration

The following are parameters that must be configured upon initial GMSU/QC installation.

26.7.1.1 Analyst Acquisition Data Path
Note:
This parameter is related to Sciex HT Acquisition Tools. If these tools will not be used, then this section may be ignored.

· This directory is the root directory of where Analyst raw data is stored. 

26.7.1.2 Assay Request File Path
Note:
This parameter is related to Sciex HT Acquisition Tools. If these tools will not be used, then this section may be ignored.

· This directory is the root directory of where Assay Request Files (or may be referred to as Compound Lists) are stored. The Compound List is essentially the text file used by Automaton™.

· This directory is generated by the GMSU/QC Setup.exe (see Section 12.5.1):

C:\GubbsInc\GMSU/QC\CompoundLists\

26.8 Hepatic Clearance Calculator

The Hepatic Clearance Calculator window contains default paths and values specific to the Hepatic Clearance Calculator Utility.
[image: image56.png]#2 Configuration Utility.

Lt | e | o] et |

Choose a category:

Caco
Fiags
Generic Viewsr
Gicbal

Hepalic Clearance Constan
Moy Method
PK Calculator

Description

Value

Default Bunching Factor Value

2

Default Chromatographic Peak Time Method

[None]

Default Data File Access Mode

2 - Open all data files within a directory

Default Directory

c:\

Default Inital Species

Rat Microsome

Default Results Plot Regression Significant Figures

s

Default Results Significant Figures

B

Default Smoothing Factor

2

Default Time Point Set

Microsome

Excel template file

Last Percent Remaining Value

150

LT or HT Data Acquisition?

i}

Sample Number to Evaluate Chromatography

B

Select Default CDS Environment

Thermo(R)

Shovs Low Throughput/High Throughput Option Windowy

No

Use 15 as Default Setting

Yes

Use UNC Path

Yes





Figure 35
Configuration Utility – Hepatic Clearance Calculator

26.8.1 Initial use parameter configuration

The following are parameters that must be configured upon initial GMSU/QC installation.

Note:
If this module will not be used, then this section may be ignored.

26.8.1.1 Default Data Directory

· This directory is the root directory of where raw data is stored. This directory is shown initially when the Browse… button is clicked in the Permeability Calculator module.

26.8.1.2 Select Default CDS Environment

Used only if users will be viewing both Thermo and Sciex data. This setting governs the order in which ‘File Type Filter’ are displayed when the Browse… button is clicked.

· If the setting is ‘Thermo(R)’, then the default choice is ‘.sld’.

· If the setting is ‘ABI(R)’, then the default choice is ‘.wiff’.

26.8.1.3 Use UNC Path

If data acquisition files (.wiff or .sld) are to be stored at the local data acquisition workstations AND the administrator wishes to configure GMSU/QC in a multi-user environment, then Use UNC Path must be set to ‘Yes’. Please see Appendix A
Thermo Environment Multiuser System Configuration, for a detailed discussion on configuring GMSU/QC in a multiuser environment.
26.9 Hepatic Clearance Calc. Expt. Constants – Existing Experiments Tab
Administrators use this window to configure species- and experiment-specific constants that are used during Hepatic Clearance value calculations. The contents of the Species dropdown boxes in the Hepatic Clearance Calculator are obtained from the Experiment list shown in Figure 36.

In this window, users configure constants, the contents of parameter labels, parameter units, Exception values, and (optional) a reference as to the origin of the constants.

As an aid to administrators when configuring new Experiments, users may copy data from one experiment to another.

To add additional experiments, use the Add Species/Type tab (see Section 26.10).
[image: image57.png]#2 Configuration Utility.

Coco
Flags i Evisting Experiments | At Species/Type Experiment Time Parts

Generic Viewsr
Bosgle Dog Hepatomte Note: Urits wil be added (o label automaticall i the Hepatic Clearance GUI
Beagle Dog Microsome Contants Labels Constants Units

Modiy Method Beagle Dog 59

P Caluiator Cynomolgus Monkey Hepatocyte|
Cynomolgus Monkey Microsome| | Protein Wt [Cell/neubval | 1 Label [Fels/incubation vol MillorCels/nl v]
Cynomolgus Monkey 59
Guinea Pig Hepatocyte.
Guinea Pig Microsome
Guinea Pig 59

Human Hepatocyte Reference: Results Labels Resaults Units
Human Microsome Intinsic Cleatance LabelUris: [[/hi/Kg

Human 59

Mini Pig Hepatogyte
Wini Pig Microsome
Mini Pig 59

Mouse Hepatogyte

Mouse Microsome. e — | ]|
Mouse 59 Intinsic Hepali Clstance. —_—

Rat Hepatogyte.
Rat Microsome
Rat 59 Eh @5

Liver Wi/Body Wit B | by |LiverwiiBody wt Ko =

ProteinWt [Cels/Liver Wt | 135 Lapel [Felis/Liver wt MilorCelis/g |

Hepalic Blood Flow 19 Label [Hepalic Blood Flow Lihikg

Predted Clearance Label Unis: [Lihi/kg

e

Fredicted Hepatic Clearance





Figure 36
Configuration Utility – Hepatic Clearance Calc. Expt. Constants – Existing Experiments Tab
26.10 Hepatic Clearance Calc. Expt. Constants – Add Species/Type Tab

Here users add species and/or types. Once added, new experiments will be generated automatically.
[image: image58.png]#2 Configuration Utility.

Caco
Fiags
Generic Viewsr

Modiy Method
PK Calculator

Beagle Dog Microsome

Beagle Dog 59

Cynomolgus Monkey Hepatocyte|

Cynomolgus Monkey Microsome|

Cynomolgus Monkey 59

Guinea Pig Hepatoyte

Guinea Pig Microsome

Guinea Pig 59

Human Hepatocyte.

Human Microsome

Human 59

Mini Pig Hepatocyte

Mini Pig Microsome’

Mini Pig 59

Mouse Hepatocyte

Mouse Microsome

Mouse 59

Rat Hepatogyte.

Rat Microsome

Rat 59

_pddspecis |

Species

Evisling Experiments

_aaype |

ad SpeciesiTyod Evpeinert Tne Parts

Beagle Dog
Cynomolgus Monkey
(Guinea Fig

Human

Mini Fig

Mouse

Rat

Adding Species and Types
wil resul in the creation
of appropriate
experiments.





Figure 37
Configuration Utility – Hepatic Clearance Calc. Expt. Constants – Add Species/Type Tab

26.11 Hepatic Clearance Calc. Expt. Constants – Experiment Time Points Tab

The Analysis Time Points window is used to create Time Point Sets that will be assigned to analysis data files. ‘Time Points’ are defined as each chromatographic peak (for HT acquisition) or each chromatographic injection (for LT acquisition).

Users click on the Edit button to configure a new Time Point Set. Users click on the Execute button to actually create the Time Point Set and the Save button to save the set. Once a Time Point Set has been created, it cannot be modified.

In order to aid internal information technology (IT) department efforts to generate internal reports or transfer data to internal laboratory information management systems (LIMS), users may label peaks not used in actual regression analysis that are needed for additional internal process control information.

[image: image59.png]#2 Configuration Utility.

Caco
Fiags
Generic Viewsr

Modiy Method
PK Calculator

Beagle Dog Microsome

Beagle Dog 59

Cynomolgus Monkey Hepatocyte|

Cynomolgus Monkey Microsome|

Cynomolgus Monkey 59

Guinea Pig Hepatoyte

Guinea Pig Microsome

Evisling Experiments

‘Add Species/Type Experiment Time Points

Time Point Sets

Microsome
so

* Optional

Guinea Pig 59

Human Hepatocyte.

Human Microsome

Human 59

Mini Pig Hepatocyte

Mini Pig Microsome’

Mini Pig 59

Mouse Hepatocyte

Mouse Microsome

Mouse 59

Rat Hepatogyte.

Rat Microsome

Rat 59





Figure 38
Configuration Utility – Hepatic Clearance Calc. Expt. Constants – Experiment Time Points Tab - Initial

[image: image60.png]#2 Configuration Utility.

Evisling Experiments ‘Add Species/Type Experiment Time Points

Beagle Dog Microsome Time Point Sets Time Points * Optional
Beagle Dog 59

Cynomolgus Monkey Hepatocyte|
Cynomolgus Monkey Microsome|
Cynomolgus Monkey 59

Guinea Pig Hepatocyte

(Guinea Pig Microsome

Guinea Fig 59

Human Hepatocyte

Human Microsome

Human 59

Mini Pig Hepatocyte Enter total number of time points:

[EZES it ity

Mouse Hepatocyte
Mouse Microsome
Mouse 59

Rat Hepatogyte.
Rat Microsome
Rat 59





Figure 39
Configuration Utility – Hepatic Clearance Calc. Expt. Constants – Experiment Time Points Tab – Edit Mode
26.12 Modify Method

Note: This window pertains only to Sciex HT Acquisition Tools.

The Modify Method window contains the contents of the LC Sync dropdown box of the Modify Methods Utility. These contents are specific to Sciex instrumentation at this point in time.

[image: image61.png]#2 Configuration Utility.

B | =

Choose a category:
Caco

Flags

Generic Viewsr
Giobal

Hepatic Clearance LC Sync Dropdown Box

B e A (Parameters lot Configurable)

P Caluiator NA

LC Sync

o Sync
Manual Sync
|1 Manif Sync





Figure 40
Configuration Utility – Modify Method

26.13 PK Calculator - Global
The PK Calculator window contains default paths and values specific to the PK Calculator Utility. Such values include default units for different PK values.

[image: image62.png]#2 Configuration Utility.

E= B = S

Choose a category:
Caco

Fiags

Generic Viewsr

Giobal

Hepalic Clearance

Hepalic Clearance Constan

oy e

Species

Description

Default Assay Set Name

Default Bunching Factor Value

1

Default Column Heading Template

Default

Default Data Directory

c:\

Default Data File Access Mode

1- Open a single data file within a directory

Default Regression

Linear

Default Results Decimal Places

2

Default Results Plot Regression Significant Figures

s

Default Results Significant Figures

B

Default Results Use Sig Figs (Yes) or Decimals (No)

Yes

Default Smoothing Factor

B

(mg=hr)/mL

mi/(Ka*hr)

Default Units: Cmax

mg/mt

Default Units: Concentration

mg/mt

Default Units: Dose

mg/Kg

Default Units: Half Life

hr

Default Weighting

1x~2

Excel template file





Figure 41
Configuration Utility – PK Calculator – Global
26.13.1 Initial use parameter configuration

The following are parameters that must be configured upon initial GMSU/QC installation.

Note:
If this module will not be used, then this section may be ignored.

26.13.1.1 Default Data Directory

This directory is the root directory of where raw data is stored. This directory is shown initially when the Browse… button is clicked in the Permeability Calculator module.

26.13.1.2 Select Default CDS Environment

Used only if users will be viewing both Thermo and Sciex data. This setting governs the order in which ‘File Type Filter’ are displayed when the Browse… button is clicked.

· If the setting is ‘Thermo(R)’, then the default choice is ‘.sld’.

· If the setting is ‘ABI(R)’, then the default choice is ‘.wiff’.

26.13.1.3 Use UNC Path

If data acquisition files (.wiff or .sld) are to be stored at the local data acquisition workstations AND the administrator wishes to configure GMSU/QC in a multi-user environment, then Use UNC Path must be set to ‘Yes’. Please see Appendix A
Thermo Environment Multiuser System Configuration, for a detailed discussion on configuring GMSU/QC in a multiuser environment.
26.14 PK Calculator – Species
Administrators configure species that will appear in the PK Calculator Species dropdown boxes.

[image: image63.png]Configuration Utility.

E= B = S

Choose a category:
Caco
Fiags

Species
Generic Viewsr
Giobal

Hepalic Clearance

Hepalic Clearance Constan

HECATEE=S Species

Dog





Figure 42
Configuration Utility – PK Calculator– Species
27 Database Table Structures and Definitions

GMSU/QC database table structures and definitions are contained in the document GMSU/QCDatabaseDefs located in C:\GubbsInc\GMSU/QC\Manuals\.
28 References

1 R. O. Cole.  Quantitation of Biofluids with High Speed HPLC/MS/MS:  Early Discovery Applications.  Invited Presentation, at “Early ADME and Toxicology in Drug Discovery: Techniques for accelerating and Optimizing Drug Candidate Selection”, Berkley, CA, October 1998.

2 Adam H. Brockman, Donna L. Hiller, and Roderic O. Cole. High Speed HPLC/MS/MS Analysis of Biological Fluids in Support of ADME Screens: A Practical Review. Current Opinion in Drug Discovery, 4(3), (2000), 432.

3 Janiszewski JS, Rogers KJ, Whalen KM, Cole MJ, Liston TE, Duchoslav E, et al: A high capacity LC/MS system for the bioanalysis of samples generated from plate-based metabolic screening. Anal Chem 2001; 73:1495-1501.

4 Z. Yan, J. Wu, L.E. Elvebak, A. Brockman: Validation of a Totally Commercially Available High Throughput -ADME System and Results for 60 Literature Compounds. Rapid Communications in Mass Spectrometry, Rapid Commun. Mass Spectrom. 2005; 19: 1191–1199.
	Appendices


Appendix A
Thermo Environment Multiuser System Configuration

It is highly recommended that, if there are any questions about the procedure of sharing a directory, local IT support is obtained since settings and permissions can vary depending on company IT policies. In fact, it is highly recommended that IT be informed that the following procedures are going to be performed.

	Table of Contents

Section
Page

	21
Introduction


22
Share the C:\Xcalibur\ Windows™ directory


23
Place the Gubbs_01.mdb in a desired location


24
Actions to take on the additional (e.g. office) workstations


25
Actions to be taken on the Thermo data acquisition workstation


26
Actions to be performed ONCE using any installation of GMSU/QC






29 Introduction

In the Thermo environment, data and processing method Windows™ directory locations are hard-coded into the Thermo sequence (.sld). For example, the location (path) of a .raw file may be recorded as C:\Xcalibur\data\02202008_122345\. Since GMSU/QC uses this path to process the .raw file, the drive letter “C:” is not helpful in a multiuser environment. If a user starts GMSU/QC on an office workstation and opens a Thermo data file (.sld file) that references ‘C:\Xcalibur\data\02202008_122345\’, GMSU/QC would fail because ‘C:\Xcalibur\data\02202008_122345\’ does not exist on the office workstation.

When the ‘Use UNC Path’ setting is changed to ‘Yes’ (see Section 6), GMSU/QC then replaces ‘C:’ with the network computer name when data references are stored in the GMSU/QC database. For example, if the data acquisition computer name is ‘Fred’, then GMSU/QC will store the .sld as the UNC path \\Fred\Xcalibur\data\02202008_122345\. In addition, all references to .raw and .pmd files made during data processing are converted to the UNC path. That way, when a user is processing data from an office workstation, GMSU/QC will open and process the data located on computer ‘Fred’.

In order to operate GMSU/QC in a multiuser environment to process Thermo data, several Windows™ directory configuration and workstation configuration actions must be performed. This document lists those actions that need to be performed.

NOTE: If this procedure is performed after data has already been processed and saved using GMSU/QC, then previously saved data can only be VIEWED using the GMSU/QC installation on the data acquisition workstation from which the data were acquired. Please note that all data may still be REPORTED (e.g. a report generated) from any GMSU/QC installation.

30 Share the C:\Xcalibur\ Windows™ directory

It is highly recommended that, if there are any questions about the procedure of sharing a directory, local IT support is obtained.

· Name the share ‘Xcalibur’

· Give share permissions to whomever you wish

31 Place the Gubbs_01.mdb in a desired location

Later steps in this process will depend on where the Gubbs_01.mdb database is placed.

In a multiuser environment, the Gubbs_01.mdb database must be accessible to all users. This can be achieved by moving the database to a network location, or by sharing the Windows directory of the Gubbs_01.mdb database current default location.

These are the options for the Gubbs_01.mdb location:

31.1 A network location

Copy and paste the C:\GubbsInc\GMSU/QC\Database\Gubbs_01.mdb to a suitable network location.

31.2 C:\GubbsInc\GMSU/QC\Database\

This is the default location. If it is desired to leave the database here, then this directory must be shared.

31.2.1 Share the C:\...\Database\ directory

· Name the share ‘GMSU/QCDatabase_01’

· Give share permissions to whomever you wish

32 Actions to take on the additional (e.g. office) workstations

The following actions must be taken on the workstation of each additional GMSU/QC installation.

NOTE: It is HIGHLY RECOMMENDED that these actions are performed under a Windows account that has Administrator permissions. If these actions are performed under a non-administrator account, then it is possible these modifications may be specific to a user profile, rather than to all users.

32.1 Modify the ODBC Datasource System DSN

The DSN on the workstation of each additional GMSU/QC installation must be modified to point to the location of the desired Gubbs_01.mdb database. Modify the DSN according to the following instructions:

32.1.1 On the additional workstation (e.g. your office workstation), open the Control Panel (Start – Control Panel)

32.1.2 Open the Administrative Tools control panel

32.1.3 Open the Datasources (ODBC) item of the control panel and activate the System DSN tab

[image: image64.png]" ODBC Data Source Administrator.

User DSN  System DSN | Fi DSN | Drivers | Tracing | Connection Pooling | About

Xireme Sample Database 2005 Microsoft Access Diver (~mdb]

the indicated data provider. A System data source is isble o al users

in ODBC System data source stores infommation about how to connect to
onthis machine, including NT services.

Cancel Aoy




32.1.4 Choose the GubbsSciex_100 DSN and click the Configure… button

[image: image65.png]ODBC Microsoft Access Setup.

Data Source Name: ST

[GmMsu/aC

System Datsbase





32.1.5 Click the Select… button

[image: image66.png]Select Database.

[Access Databases ('m +





· Navigate to the directory that contains the location of the GMSU/QC_01.mdb as performed in Section 3.

· The Network… button will need to be used

· Select the Gubbs_01.mdb database

· Click OK

32.1.6 Exit the control panel

The configuration is finished. Click OK in the rest of the control panel windows to exit the control panel.

33 Actions to be taken on the Thermo data acquisition workstation

33.1 If Section 3.2 was executed (the Gubbs_01.mdb database was not moved from its default location), then no changes need to be performed on the data acquisition workstation.

33.2 If Section 3.1 was executed (the Gubbs_01.mdb database was copied and pasted to a network location), then repeat the steps listed in Section 4, but perform on the data acquisition workstation instead.

34 Actions to be performed ONCE using any installation of GMSU/QC
Once the workstation configuration steps have been completed, then GMSU/QC must be configured to operate in a multiuser environment. Each GMSU/QC module that will be used in a multiuser environment must be configured.

Since all GMSU/QC installations are now pointing to the same database, this section needs to be performed on once for each module. The changes will affect all GMSU/QC installations.

This process will be described using the Generic Chromatographic Viewer as an example.

34.1 Start GMSU/QC and open the Configuration Utility

[image: image67.png]3 GMsU Console

Gubbs® Mass Spec Utilities - distributed as Thermo® QuickCalc® | About..|

_tib.. |

- Home Sciex® HT Acquisition Tools

| Hepatic Clearance Calculator
Permeability Calculator

Generic Chromatographic Viewer
PK Calculator

Configuration Utility

Exclusively Disrbuted 35 DuickCalo® By: ‘Watkstation Prerequisies Installedt A product o Gubbs Inc

Thermo SLVRRY
SCIENTIFIC





34.2 Select the Generic Viewer category

Scroll through the Global list until the ‘Use UNC Path’ item is found

[image: image68.png]#2 Configuration Utility.

Giobal
Hepalic Clearance

Hepalic Clearance Constan
Moy Method

PK Calculator

Pemeabily

Description

Default Results Decimal Places

B

Default Results Plot Regression Significant Figures

s

Default Results Significant Figures

B

Default Results Use Sig Figs (Yes) or Decimals (No)

Yes

Default Smoothing Factor

7

Default Units: Concentration

mg/me

Excel template file

C:\GubbsInc\GMSU \PostReportMacros\BMS\BMSTemplate.xis

Flag Criteria (%) - Analyte Retention Time

10

Flag Criteria (%) - Int Std Peak Area

10

Flag Criteria (%) - Int Std Retention Time

10

Flag Criteria (%) - QC Peak Area

20

If HT Acquisition, default chrom view is T

LT or HT Data Acquisition?

i}

Sample Number to Evaluate Chromatography

4

Select Default CDS Environment

Thermo(R)

Show biinking flag labels

Yes

Shovs Low Throughput/High Throughput Option Windowy

Yes

Use UNC Path No

<

Use IS as Default Setting

Yes

3





34.3 Change the value to ‘Yes’

· Select the ‘Use UNC Path’ item

· Click on the Edit button

· Select ‘Yes’ as the New Value

· Click on the Save button

[image: image69.png]Global
Description  [ooleUNE ]
oavae o
newvaive [,
| -





34.4 Repeat the process (if applicable)

Repeat Section 6 for each module that will be used in a multiuser environment:

· Generic Viewer

· Hepatic Clearance

· PK Calculator

· Permeability

Version 7.0 (December 27, 2008)
C:\GubbsIncApps\GMSU\Administration\Manuals\Admin\GMSU_AdminManual_0007.doc

[image: image70.jpg][image: image71.jpg]%VKK?



